## CS-Digital Cell Analyzer
CS File of Digital Cell Analyzer – Research & Development (R&D)
Requirement | Incepta Specification / Description | Vendor Feedback
Vendor / Product Information
Manufacturer | NA
Local Agent | NA
Origin | NA
Model | NA
1. General Information
Description | Cell counter system designed for rapid and accurate determination of cell concentration, viability, and morphology using image-based analysis technology. Supports analysis of suspension and adherent cell cultures with minimal sample volume. Provides automated data acquisition, result analysis, graphical representation, data storage, and export functionalities suitable for research and laboratory applications.
Sample Analysis Capacity | Capable of analyzing viable and total cell concentration, viability, and cell morphology for multiple samples.
Data Storage & Analysis | Data storage and analysis through PC software
Utilization – Hours per day | 24 hours per day
Utilization – Days per week | 7 days per week
Room Temperature | NMT 27°C
Relative Humidity | 20–80%
Zoning | Grade C area or Laboratory
Pressure | 10–15 Pa
2. Functional / Operational Requirements – Design Parameters
Capable Samples | Mammalian and other eukaryotic cells
Cell Culture Type | Suspension or Adherent culture
Capable Cell Size Range | 5–180 µm
Cell Concentration Range | 1×10⁴ – 3×10⁷ Cells/mL
Analysis Method | Image-based automated cell analysis
Sample Loading Volume | 10–50 µL
Viability Method | Trypan Blue exclusion
Slide Type | Disposable 4-chamber slide or more chambers
Measurement Parameters | Total cell count, viable cell count, dead cell count, viability, cell diameter, roundness, and aggregation (density) analysis
Graphical Analysis | Diameter histogram, aggregation histogram, cultivation time chart, and growth curve analysis
Analysis Time | Approximately 10–60 seconds per sample
Data Storage & Analysis | PC-based data storage, analysis, retrieval, and export facility
Data Export Format | Compatible with PDF, XLS, and image export formats
Imaging System | Integrated automated imaging system with autofocus capability
Software Features | Histogram analysis, growth curve analysis, overlay comparison, and historical data management
Compatible Operating System | Windows-based operating system
Dimensions | Benchtop-sized compact design suitable for laboratory use
Connectivity | USB / LAN
2.2 Controlling System
PC | External PC
System Control | System control software
Imaging System Control | Automated image capture and processing
Startup / Shutdown | Manual startup / shutdown
Operation Mode | Switch / Push button operated
Display | LED display (Optional)
2.3 Measuring / Calibration Parameters
Cell Density Calibration | Required – based on equipment specification
Cell Viability Calibration | Required – based on equipment specification
Cell Diameter Calibration | Required – based on equipment specification
2.4 Safety & Alarms
Standard Alarms / Error Indication | Equipment should show standard alarms or errors (e.g. successful run completion, run failure, equipment malfunction etc.)
Parameter Deviation Alarm | Alarm during any deviation from set parameter
2.5 Power Failure & Recovery
Power Recovery | Automatic power recovery system at previous set point after a power failure
2.6 Material of Construction (MOC)
Material Compatibility | Compatible material for cell culture process and cleaning
2.7 Ink Jet / USB / SD Card Supported Print Out
Print Out Option | N/A
3. Utility Services
Voltage | 100V to 240V, 5W
Frequency | 50–60 Hz
Uninterrupted Power Supply | Required
4. Validation Aspects / Regulatory Requirements
IQ-OQ Installation | By suppliers
IQ-OQ Documents | Required
List of Components | Required
Calibration Certificate | Required
Cleaning Requirements | Smooth and easily cleanable by using usual cleaning agent
21 CFR Part 11 / Data Integrity | N/A
Document Language | Documents shall be in English version and each document shall be uniquely traceable
5. Maintenance / Calibration
Calibration During Installation | Manufacturer or vendor will bring calibrated instruments and corresponding valid calibration certificate for qualification of equipment (if applicable); otherwise calibration will be performed during installation.
Technical Support | Technical support during installation, operation and emergency purpose will be provided by supplier.